DEFENCE RESEARCH & ELECTRONICS TESTING LABORATORY

NOTICE INVITING TENDER: DETL/PROC/LAB-PS/2026/15

TENDER FOR: Supply of High-Precision Programmable Laboratory DC Power Supplies (0-60V, 0-30A, 900W)

SECTION 3: TECHNICAL SPECIFICATIONS FOR LABORATORY DC POWER SUPPLY

1. OUTPUT RATINGS & PERFORMANCE:

1.1 Output Voltage: Continuously programmable from 0 to 60 Volts DC.

1.2 Output Current: Continuously programmable from 0 to 30 Amperes DC.

1.3 Output Power: Total continuous output power shall be minimum 900 Watts with autoranging V/I envelope.

1.4 Line Regulation: <= 0.01% + 2 mV for voltage; <= 0.01% + 250 uA for current.

1.5 Load Regulation: <= 0.01% + 2 mV for voltage; <= 0.01% + 250 uA for current.

1.6 Voltage Ripple and Noise (20 Hz to 20 MHz): Not exceeding 2 mVrms / 20 mV peak-to-peak.

1.7 Current Ripple: Less than 5 mArms.

2. ACCURACY & PROGRAMMING RESOLUTION:

2.1 Programming Resolution: Voltage resolution <= 1 mV; Current resolution <= 1 mA.

2.2 Readback Accuracy: Voltage +/- (0.05% + 10 mV); Current +/- (0.1% + 10 mA).

2.3 Transient Response Time: Less than 50 microseconds for output recovery to within 15 mV following a 50% to 100% load step.

3. OPERATING MODES & PROTECTIONS:

3.1 Seamless automatic crossover between Constant Voltage (CV) and Constant Current (CC) modes.

3.2 Comprehensive hardware and software protection: Over Voltage Protection (OVP), Over Current Protection (OCP), Over Temperature Protection (OTP), and Reverse Voltage Polarity Protection.

3.3 Output enable/disable electronic soft switch with isolated floating output terminals.

4. USER INTERFACE & REMOTE CONNECTIVITY:

4.1 Minimum 4.3-inch high-resolution color TFT display showing simultaneous real-time voltage, current, power, and operational mode.

4.2 Standard remote programming interfaces: USB (USBTMC), RS-232, and LAN (LXI Core 2011 compliant) with SCPI command set.

4.3 Analog programming and monitoring interface (0-5V / 0-10V control inputs).

5. ELECTRICAL & ENVIRONMENTAL SPECIFICATIONS:

5.1 Universal AC input: 100V to 240V AC (+/- 10%), single phase, 47-63 Hz with active Power Factor Correction (PFC >= 0.98).

5.2 Operating temperature: 0°C to 40°C, Relative humidity: up to 80% non-condensing.

5.3 Cooling: Forced air cooling with low-noise temperature-controlled smart fan.

5.4 Chassis earthing terminal complying with IS 3043 Code of Practice for Earthing.

6. SAFETY & QUALITY STANDARDS:

6.1 Safety conformity: Conformance to IS/IEC 61010-1 Safety Requirements for Electrical Equipment for Measurement, Control, and Laboratory Use.

6.2 Electromagnetic Compatibility (EMC): Conformance to IS/IEC 61326-1 for laboratory equipment and IS/IEC 61000-3-2 for AC harmonic emissions.

6.3 Calibration: Valid calibration certificate from an NABL accredited laboratory (ISO/IEC 17025) with reported measurement uncertainties.

6.4 3 Years Comprehensive OEM warranty including parts and calibration support.